## Slide 2: The Theory of Near Light Scatter
Scattered light has components due to diffraction, refraction and reflection
For low detection angles, scattered light is dominated by the diffraction and refracted components
The particulate reflective properties and colour have minimal effect on instruments response
Diffracted and refracted intensities are small

## Slide 3: Microdust Photometer layout

## Slide 4: Clean air condition
Infrared beam is focused onto the light stop
No forward scattering effect
Zero signal intensity seen by detector

## Slide 5: Particulate introduced to sample volume
Beam is scattered due to the effects of diffraction, refraction and reflection
Intensity of Infra red signal is measured by detector

## Slide 6
Concept

## Slide 7
ADR 1500

## Slide 8
PDR 1100

## Slide 9: ADR 1500

## Slide 10: Principle of Operation
What airborne pollutants can be measured
Data Ram responds to all particulate matter that is airborne. It is however optically configured to produce optimal response to particles in the range 0.1 – 10.0 µm, achieving a correlation to gravimetric measurements.
Output
Data Ram produces an analog 4-20 mA powered loop and a 0-5 V output signal. It also has a selectable Alarm digital output.
Software
Data Ram is supported by pDR-COM custom software, however this is aimed at short-term personal monitoring and is not suitable for continuous long term fixed-point operation. The PDR-1100 output signal is normally fed into the “Backbone” system of the mine, all reporting and data logging is configured by the end user.

## Slide 11: Applicability (surface, underground, plants)
PDR1100 was developed specifically for underground use where dust loading is typically high and consequential contamination of sensing devices and other components are a reality. PDR1100 does therefore not employ a pump. Air is induced at low velocity which entrains the smaller particles and reduces contamination.

## Slide 12: Ex certification
SANS 60079-: “Electrical apparatus for explosive gas atmospheres”;
  Part-0: 2009 “General Requirements”; and
  Part-11: 2007 “Equipment protected by intrinsic safety ‘I”
Approved explosion protection rating: Ex ib I
Inspection Authority Certificate: MASC M/10-259X

## Slide 13: Installation
PDR1100 is supplied with two steel suspension Eye Bolts. A directional arrow indicates the mounting position in relation to the mine airflow. The electrical installation is a simple 3-wire connection. When power is applied, the instrument switches on and starts transmitting automatically.

## Slide 14: “Backbone Requirements”
Power supply: 12-24 V DC (“IS” for Fiery Mines)
Output: 4-20 mA or 0-5 V

## Slide 15: ADR 1500 SOLAR PANEL

## Slide 16: Maintenance requirements
Annual Calibration – This is carried out at our Roodepoort facility. The process includes cleaning and inspection of the Sensor Lenses and other internals. The zero process is done on a certified Clean Air Bench which uses Hepa Filters. The Span calibration is carried out in a Ventilated Dust Chamber using Arizona Road Dust - ISO SAE Fine.
Field Calibration – This is essentially a zeroing function which resets the zero point above the background level of accumulated dust in the measuring chamber. The PDR1100 is equipped with ports which aspirated to remove accumulated dust from the Measuring Chamber.

## Slide 17: Supplier support network
Locally manufactured or imported
Data Ram is manufactured by Thermo Fisher Scientific in Franklin, Massachusetts USA. The outer housing and ancillaries are manufactured locally.
Our HQ
AMS Haden Instrument Service are based in Allen’s Nek, Roodepoort. All sales and technical support is provided from this centre.

## Slide 18: OEM Training and support
AMS Haden provide on-site user training on all aspects of PDR deployment. The company also provides a comprehensive repair and calibration service from our Roodepoort facility.

## Slide 19
PDM 3700

## Slide 20
The End